Combi-PAL Installation Qualification

Execution Approval Document

The signatures below represent approval of all interested parties for the execution of the LEAP Technologies Combi-PAL Installation Qualification Protocol.

Company Name:____________________________________________________

Name:			Signature:			Title:			Date:

## LEAP Technologies Representative

Name:			Signature:			Title:			Date:

### Signature Page

The signatures below represent all those individuals who may have executed or assisted in the execution of the LEAP Combi-Pal Installation Qualification Protocol.

Company Name:____________________________________________________

Name:			Signature/Title:		Initials		Date:

## LEAP Technologies Representative:

Name:			Signature/Title:		Initials		Date:

###### Table of Contents

- Table of Contents

- Revision History

- PAL System Identification

- Scope / Purpose

- Documentation

###### Procedure

- Laboratory Requirements

- Temperature

- Space Requirement

- Power Requirement

- Safety

- Accessories

- Packing List

- Installation

- Object Alignment

- Functionality Test

- Training /Familiarization

- Enclosures

- Certification

- Non-Conformance Record

II.	Revision History

| Revision | Date | Author | Reason for Changes |
| --- | --- | --- | --- |
| Rev A | June 10, 2001 | Ingo Christ LEAP Technologies | Initial Protocol |
| Rev B | June 3, 2002 | Ingo Christ LEAP Technologies | Added customer acceptance sign-off to IQ certificate. Added N/A option to inventory list. |
| 1.0 | Jan. 22, 2004 | Eric Wethington and Michael Horton LEAP Technologies, Inc. | Format to conform to LEAP Quality System |

###### Combi-PAL System Identification

| Serial Number: |  |
| --- | --- |
| Customer: |  |
| Address: |  |
| Building/Location: |  |
| Main Contact: |  |
| Distributor: |  |
| Cruise Control: |  |
| Date Install: |  |
| Warranty until: |  |

###### Scope / Purpose

The objective of this document is to define the Installation Qualification requirements and acceptance criteria for the PAL System.

The Installation Qualification will document the completeness of shipment the location and the proper mechanical functionality of the instrument. It will furthermore verify the correct identification of the system.

Successful completion of this document will provide a high degree of assurance that the PAL system has been installed to LEAP Technologies specification.

Only certified personnel from LEAP Technologies can perform the Installation Qualification. A Certification letter of the acting service engineer is attached to this document.

Prior to the Installation Qualification the customer is required to read and approve the Procedure. All specified equipment and conditions must be met for a successful validation.

###### Documentation

All qualification work will be performed as defined in the approved protocol and all the documentation will be completed concurrently with the execution of the qualification studies. Data entries must be completed during the qualification testing. Any deviation or exceptional condition observed during the execution of the protocol must be documented in the attached “IQ Non-Conformance Record”.

If any point or section is not applicable for the Specific PAL System it shall be marked with “N/A” (= not applicable).

All required documentation must be filed together in one location.

The documentation will be reviewed and approved by the lab management who initially approved the qualification studies.

A sign-off section is provided on each page of this document

Note: This page must be signed by both parties to confirm this agreement.

___________________________________________		___________________________________________

LEAP Representative		Date			Company Representative	Date

###### Procedure

The Instrument Qualification is split in three sections. This is an overview of all action, which will take place.

### Correct installation of hardware

Only the certified LEAP personnel can perform the installation procedure. LEAP personnel has to follow explicitly the installation procedure recommended by the vendor. Only described variations due to space limitation can be taken. All necessary environmental conditions must be met. The customer must supply proper power condition, high-grade chemicals and gases.

### Locate and document item of the product

All items of the shipment need to be checked against the P.O. A copy of the P.O. needs to be attached to this document. All serial numbers, Firmware Versions and Software Versions must be documented. Note all changes of the standard stetting of the system and the customer must sign off those changes.

### Proper mechanical functionality

After the installation of all the objects attached to the system, they need to be checked via the handheld remote and the check position feature. Several positions within the tray need to be checked. The safety guard must be attached to the system.

###### Laboratory Requirements

- Temperature

Verify that the environmental requirements for safe operations are met for proper installation and use.

Temperature 10 C to 30 C			   Actual Temperature:_______C   Pass (  )

_______F    Fail  (  )

Indoor and protected use only			   			     Pass  (   )  Fail  (  )

- Space Requirements

The dimensions and weight of the PAL System allow the system to be placed on almost any laboratory bench. It is important to verify that the system can be placed in a horizontal position and that the bench is designed to carry the weight. Add a few centimeters to the minimum measurement around the system to allow space for cables and tubing.

GC PAL, DI PAL, Combi-PAL:	Width:	830 mm

Depth:	390 mm

Height:	650 mm

- Power Requirements

Verify that the power supply is designed for the correct operating voltage of:

97.5 – 264 VAC

- Safety Measure / Precautions

The safety measures are outlined in the User Manual, page 2 and in any “Addendum to the User Manual”, e.g. the Fiber Bake-out Station, if applicable.

The Safety Guard must be installed.  If the user declines the installation, it shall be remarked in the “Non-Conformation Record”.

- Accessories

The customer is responsible for a proper maintained and functional chromatographic system. All supplies necessary to run or to maintain the instrument need to be available, such as head space and direct liquid injection syringes, high purity gases for carrier gas and detector, wash solvents etc. LEAP will supply 250 L liquid syringes and certified thermocouples.

###### Packing List

This is a complete packing list with all parts necessary for a Combi-PAL. Some parts might not be ordered from the customer. Use the Purchase Order for a cross check.

Check if available 	ordered	N/A		Description

 		 				1 X/Y crossrail unit;

 						1 Injection unit (Z unit);

	 					1 Four solvent wash station;

	 					1 Keypad controller, with cable;

	 					1 Keypad bracket;

			 			1 liquid Syringe adaptor;

 						1 10 uL syringes;

 						1 2.5 mL HS syringes;

 						1 Heated syringe adapter;

 						1 Interface cable, PAL to analyzer;

  						1 RS232 cable;

 						2 Tray holders;

 						2 32 position trays;

 						1 98 position tray;

 						1 Power supply, with cord and cable;

 						1 Operator manual;

 						1 Acrylic guard;

 						2 Support feet;

 						1 Agitator;

 						2 Agitator cables, 1 long, 1 short;

 						6 Agitator inserts for 10 mL vials;

 						1 20 mm crimper;

 						1 Start up kit;

 						25 20 mL vials;

 						25 10 mL vials;

 						50 Caps and septa;

  						1 Regulator;

 						3 Torx Drivers, sizes 10, 20 and 25;

 						1 Hex wrench, 1.5 mm.

_____	______________________	__________________________

Signature Technician			Signature Customer

###### Installation

#### Description									 	 pass, fail, N/A

- Verify proper voltage and load rating of the PAL Prep and Load System and the	______ 	                  wall receptacle.

- The top of the GC should not have any objects mounted on. If so remove them.	______

- Install the two brackets to the GC. Depending what brand and model the PAL 		______	     auto-sampler is assembled onto, the bracket screw / bolt on to the side panel or the 			       top of back panel.

- Place the X-Y crossrail on top of mounting legs into the groove of the X crossrail 	______		       and tighten them with the Torx screws. Decide in advance on which side the 			            agitator will be attached to. Make sure that you can reach position 3 or 6 if 			           mounted on the right hand or position 1 and 4 if mounted on the left.

- Mount the injection head to the Y crossrail utilizing the 3 screws attached and 		______		     plug the ribbon cable connection to the injection head.

- Install the guard assembly.								______

- Add the Keypad terminal to the left or right side of the crossrail depending on 		______		        the individual configuration.

- Add the Agitator to the left or right side depending on lab space or preference 		______	         of customer.

- Add the wash station to the X crossrail groove and tighten the Torx screws. 		______		      The position of the wash station to be determined by the individual configuration. 		           Position the wash station close to the between the trays and the inlet.

- Fill the Wash station with the appropriate solvents.					______

- Add Tray Holders or Stacks by attaching to the X crossrail groove and tighten the 	______		     Torx screws. The position of the Holders or Stacks to be determined by the 			        Individual configuration.

- Mount the liquid syringe holder with appropriate syringe to the Injection Head.	______

- The supply gas pressure for the Headspace flushing line should be less that 25 psig. 	______	  Please verify for leaks. If only higher pressure gas lines are available add a needle 			    valve or second regulator. Verify and correct for leaks.

- Make all electrical connections (Agitator cable, controller cable, power module 	______		   cable and others) from the accessories to the X crossrail.

- Plug in the power cable to the power module and the Cooled Tray if available.		______

- Turn on the power to the PAL Prep and Load system.					______

- Turn on the power to the Peltier Cooled Tray if applicable. Set a Temp of 10 C.	______

Verify that Peltier Cooled Tray cools to set temperature.

###### Object Alignment

#### Description									 	 pass, fail, N/A

## Locate the object reference position of the Trayholders or Stacks by locating 		______        the hole in the center of the Trayholders or Stacks (top drawer must be opened          Manually to access).

- Adjust the X, Y and Z positions and verify by using the F1 Function key 		______ “Check Positions” for each Trayholder or Stack.

- Locate the Wash Station’s X, Y and Z positions for Wash1 and Wash2. Verify by 	______     using the F1 Function key  “Check Positions”.

- Use the Change Syringe Function, insert the syringe into the syringe holder on the    	______ Injection Head.

- Prime the Syringe utilizing the “Clean Syr” with the “Wash1 and/or the Wash2”       	______ Solvents.

- Verify the Tray sizes to be used with the Trayholders or Stacks.				______

- Locate the Tray’s  X, Y and Z positions for Sample positions. Verify by using 		______        the F1 Function key  “Check Positions”. Repeat for each Tray.

- Locate all inlets X, Y and Z positions. Verify by using F1 key and penetrate the 	______    septa by checking the 40 mm injection depth.

- Verify all other syringe types if applicable. Verify the proper detection of each           	______ syringe.

- Check the Remote Start Cable from the PAL to the GC. All other device will be 	______ started from the GC. If PAL is installed to existing systems it might be necessary to reconfigure the starting process (e.g. HP 5890 GC with 5972 MS). A detailed list 	         of attached devices must be send to the installing service engineer in advance. 	   Often a specialized cable configuration is required. To avoid any unnecessary 		  delays this information is crucial for a successful and timely IQ.

- Confirm the Sync Signals and the injected Output Signals.				______

- Install the Pal Loader program and Cruise Control appropriate Drivers for 	   	______ Chemstation or Xcalibur (if available) unto the PC.

- Run the PAL Loader program and make a Backup of the PAL system. Keep a 		______      copy for the customer and send another copy to LEAP Technologies.

- Reset the PAL Prep and Load system after completion.					______

- Disconnect the serial cable from the portable PC.					______

###### Functionality test

#### Description									 	 pass, fail, N/A

- Add 2mL vials to the VT98 tray and check following positions via the 		         	______ handheld remote:  Setup\Objects\Trays\Select Tray\F3 and go to Position 		          1, 14, and 98. A 2 mm distance to the vial is acceptable.

- Add crimped 20 mL vials and check following positions on Tray VT32-20 via the     	______ handheld remote: Go to Position 1, 8, and 32. Add crimped 20 mL vials to 	            position 1, 8, and 32 in order to see if the arm passes the height of those vials.

- Change the syringe a check their automatic detection. Use the Menu/F1 function 	______         to exchange the syringe. Make sure to check no syringe as well.

- Attach the headspace syringe and holder. Check the temperature of the heated 	          	______  agitator and the HS syringe via the utility/tray/agitator menu and the utility/syringe 	   menu.

- Verify the sync and start signals. Start a new job and program a short sequence 	______         on the PAL. Use following variables:

| Combi-PAL conditions |  |
| --- | --- |
| Cycle | Liquid Injection |
| Syringe | 10 uL |
| Fill Speed | 5 uL/s |
| Inject Speed | 100 uL/s |
| Pre Inject Delay | 0 ms |
| Post Inject Delay | 0 ms |
| Pre Clean Syringe w Solvent One | 1 |
| Pre Clean Syringe w Solvent Two | 0 |
| Post Clean Syringe w Solvent One | 0 |
| Post Clean Syringe w Solvent Two | 0 |

Program an isothermal method with an appropriate split, oven temperature, and

runtime.  Start a sequence table with two samples and make your data analysis software

wait for a manual injection. Verify that the sampler waits for a ready signal and starts 	        the GC at the appropriate time. Review both data files and check for appropriate runtime.

- Check the Pal Loader Software for functionality and take a backup of the current    	______ firmware.

- If purchased, check the Cruise Control or Cycle Composer for functionality. Run a 	______ short sequence with Cycle Composer and make sure that the system starts the external systems. Reload the Pal Configuration and make the Customer aware of the option 	 Format Sample List under File/Options/Format Sample List. Depending on which 	 macros are activated different objects in the sample list are necessary.

###### Training / Familiarization

This is essential for a new installation of the PAL system. Only trained users should operate it.  It is best if one user assumes responsibility for any questions regarding maintenance, manuals and technical expertise.  The initial training is performed by the LEAP engineer who performed the original installation.  He should be the main point of contact for any application and usability questions.

The training provided by the third-party validation company will focus on method validation and how to keep the system in compliance. Other optional topics are usability, basic method development, basic programming via the handheld remote, starting signal configuration and the realignment and backup of the system. The importance of regular backups and routine maintenance should be emphasized.

_____________________	________________________	______________

Attendee			Company			Phone Number

_____________________	________________________	______________

Attendee			Company			Phone Number

_____________________	________________________	______________

Attendee			Company			Phone Number

_____________________	________________________	______________

Attendee			Company			Phone Number

###### Enclosures

The following documents are filed and stored in the correspondence LOG Book for the PAL System.

Copy of P.O. (for new units)

Delivery Note or Packing List (for new units)

Certificate of Test Conformation

Non-Conformance Record

Certification letter of service engineer

Certificate of Temperature Measuring Device

Other documentation

###### Certification

Installation Qualification

# CERTIFICATE OF CONFORMANCE

## Combi-PAL

Serial No.________________________

#### Has successfully passed all installation qualifying specifications

Date________________________

______________________________________    ____________

Certification Engineer                                           Date

______________________________________    ____________

Customer Acceptance	                                     Date

###### Non-Conformance Record

| Ref # | Detailed Description of Non- Conformance | Action taken | Accepted changes Initial and Date by both parties |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| 4 |  |  |  |
| 5 |  |  |  |
| 6 |  |  |  |
| 7 |  |  |  |
| 8 |  |  |  |
| 9 |  |  |  |

# LEAP Combi-PAL IQ Validation Document

Execution Final Approval

The signatures below represent approval of all interested parties for the final acceptance of the completed LEAP Combi-PAL Installation Qualification Protocol.

Company Name:____________________________________________________

Name:			Signature:			Title:			Date:

## LEAP Technologies Representative

Name:			Signature:			Title:			Date: