# HTS-PAL Installation Qualification

Execution Approval Document

The signatures below represent approval of all interested parties for the execution of the LEAP Technologies HTS-PAL Auto-sampler IQ Validation Protocol.

Company Name____________________________________________________:

Name:			Signature:			Title:			Date:

## LEAP Technologies Representative

Name:			Signature:			Title:			Date:

### Signature Page

The signatures below represent all those individuals who may have executed or assisted in the execution of the LEAP HTS-Pal Installation Qualification Protocol.

Company Name____________________________________________________:

Name:			Signature/Title:		Initials		Date:

## LEAP Technologies Representative:

Name:			Signature/Title:		Initials		Date:

###### Table of Contents

- Table of Contents

- Revision History

- PAL System Identification

- Scope / Purpose

- Documentation

###### Procedure

- Laboratory Requirements

- Temperature

- Space Requirement

- Power Requirement

- Safety

- Accessories

- Packing List

- Installation

- Object Alignment

- Functionality Test

- Training /Familiarization

- Enclosures

- Certification

- Non-Conformance Record

II.	Revision History

| Revision | Date | Author | Reason for Changes |
| --- | --- | --- | --- |
| 1.0 | April 19, 2004 | Eric Wethington LEAP Technologies | Reformatted to conform to LEAP quality system |

###### III.	HTS-PAL System Identification

| Serial Number: |  |
| --- | --- |
| Customer: |  |
| Address: |  |
| Building/Location: |  |
| Main Contact: |  |
| Distributor: |  |
| Cruise Control: |  |
| Date Install: |  |
| Warranty until: |  |

- Scope / Purpose

The objective of this document is to define the Installation Qualification requirements and acceptance criteria for the HTS-PAL System.

The Installation Qualification will document the completeness of shipment the location and the proper mechanical functionality of the instrument. It will furthermore verify the correct identification of the system.

Successful completion of this document will provide a high degree of assurance that the HTS-PAL system has been installed to LEAP Technologies specification.

Only certified personnel from LEAP Technologies can perform the Installation Qualification. A Certification letter of the acting service engineer is attached to this document.

Prior to the Installation Qualification the customer is required to read and approve the Procedure. All specified equipment and conditions must be met for a successful validation.

## Documentation

All qualification work will be performed as defined in the approved document (protocol), and all the documentation will be completed concurrently with the execution of the qualification studies. Data entries must be completed during the qualification testing.

Any deviation or exceptional case observed during the process of the OQ must be documented in the attached “OQ Non Conformance Record”.

If any point or section is not applicable for the specific HTS-PAL Auto-sampler System to be tested, it shall be marked with N/A (= not applicable).

All required documentation must be filed together in one location.  To ease this process, LEAP provides the customer with a System Log book.

The documentation will be reviewed and approved by the lab management who initially approved the qualification studies.

A sign-off section is provided on each page of this document

Note: This page must be signed by both parties to confirm this agreement.

_______________________________________       ___________________________________________

LEAP Representative	Date			Company Representative	Date

- Procedure

The Instrument Qualification is split in three sections. This is an overview of all action, which will take place.

### Correct installation of hardware

Only the certified LEAP personnel can perform the installation procedure. LEAP personnel has to follow explicitly the installation procedure recommended by the vendor. Only described variations due to space limitation can be taken. All necessary environmental conditions must be met. The customer must supply proper power condition, high-grade chemicals and gases.

### Locate and document item of the product

All items of the shipment need to be checked against the P.O. A copy of the P.O. needs to be attached to this document. All serial numbers, Firmware Versions and Software Versions must be documented. Note all changes of the standard stetting of the system and the customer must sign off those changes.

### Proper mechanical functionality

After the installation of all the objects attached to the system, they need to be checked via the handheld remote and the check position feature. Several positions within the tray need to be checked. The safety guard must be attached to the system.

## Laboratory Requirements

- Temperature

Verify that the environmental requirements for safe operations are met for proper installation and use.

Temperature 10 C to 30 C			Actual Room Temperature:	_________

Pass (  ) Fail (  )

Relative Humidity < 80%, non condensing

Indoor and protected use only.

- Space Requirements

The dimensions and weight of the PAL System allow the system to be placed on almost any laboratory bench. It is important to verify that the system can be placed in a horizontal position and that the bench is designed to carry the weight. Add a few centimeters to the minimum measurement around the system to allow space for cables and tubing.

HTS-PAL

Width:	830 mm

Depth:	385 mm

Height:	648 mm

- Power Requirements

Verify that the power supply is designed for the correct operating voltage of:

97.5 – 264 VAC, 45 – 66 Hz.

- Safety Measure / Precautions

The safety measures are outlined in the User Manual, page 2 and in any “Addendum to the User Manual”, e.g. cooled stack, if applicable.

Note: The Safety Guard must be installed. If the user declines the installation, it shall be remarked in the “Non-Conformation Record”.

- Accessories

The customer is responsible for a proper maintained and functional chromatographic system. All supplies necessary to run or to maintain the instrument need to be available, such as high purity solvents and detector, wash solvents, deionized water etc.

LEAP will supply syringes and thermocouples.

- Packing List

This is a complete packing list with all parts necessary for a HTS PAL. Some parts might not be ordered from the customer. Use the Purchase Order for a cross check.

Check if available 	ordered	Description

 				1 X/Y cross-rail unit;

 				1 Injection unit (Z unit)

	 			1 Fast Wash Station Valve Assembly

	 			1 Wash station bracket

	 			2 500 uL Bottles

	 			1 tubing kit, FWS to Bottles

	 			1 Cable, PAL to FWS

	 			1 Drainage tube for FWS

	 			1 Drainage tube for valve

	 			3 Valve insert sleeves

	 			3 Screws for attaching the Injection Head to the Y Axis

	 			3 Extra screws for attaching the Injection head to the Y						 Axis

	 			1 Keypad controller, with cable

	 			1 Keypad bracket

	 			1 Plunger holder

	 			1 Syringe adaptor

	 			1 100ul syringes (Default)

	 			1 Valve Drive, with 6 port, 2 position SS valve

	 			1 Interface cable, PAL to analyzer

	 			1 Null modem cable

	 			1 Power supply, with cord and cable

	 			1 Operator Manual

	 			1 Acrylic guard

	 			2 Support feet, standard configuration with 2 x 215 mm						 legs attached to 2 x 280 mm feet

	 			2 auxiliary cables, 1 long, 1 short

	 			3 Torx Drivers, sizes 10, 20 and 25

	 			1 hex wrench, 2.0 mm

	 			1 Manual

	 			1 Loader

_________________			________________

Signature Technician			Signature Customer

- Installation

#### Description									 	 pass, fail, N/A

- Inspect container for visible damage.							______

- Verify proper voltage and load rating of the PAL Prep and Load System and the	______ wall receptacle.

- The bench space for the PAL Prep and Load System must be flat and even. 		______

- Place the X-Y cross-rail on top of mounting legs into the groove of the X cross-rail 	______ and tighten them with the Torx screws.

- Set the two brackets beside to the LC. Depending what brand and model the PAL 	______ autosampler is assembled onto, the bracket screw / bolt on to the side panel or the 		      top of  back panel.

- Mount the injection head to the Y cross-rail utilizing the 3 screws attached and plug 	______ the ribbon cable connection to the injection head.

- Install the guard assembly.								______

- Add the Keypad terminal to the left or right side of the cross-rail depending on the 	______ individual configuration.

- Add the Valve by attaching it to the X cross-rail groove and tighten the Torx screws.	______ The position of the Valve to be determined by the individual configuration. Connect 	 sample loop, waste, carrier/mobile phase and column lines to Valve body.

- Add the fast wash station to the X cross-rail groove and tighten the Torx screws. 	______

- The position of the fast wash station to be determined by the individual configuration.	______ Position the fast wash station close to Valve assembly.

- Fill the Solvent Reservoirs with the appropriate solvents for the as designated by 	______

- Wash1 and Wash2 and install unto the solvent rack of the fast wash station.		______

- Add Tray Holders or Stacks by attaching to the X cross-rail groove and tighten the 	______ Torx screws. The position of the Holders or Stacks to be determined by the 	        Individual configuration.

- Mount the liquid syringe holder with appropriate syringe to the Injection Head.	______

- Make all electrical connections (Valve cable, controller cable, fast wash cable, power 	______ module cable) from the accessories to the X cross-rail.

- Plug in the power cable to the power module and the Cool stack if available.		______

- Turn on the power to the PAL Prep and Load system.					______

- The PAL Prep and Load system powers on.						______

- Turn on the power to the Cool Stack if available. Set a Temp of 10 ºC.			______

Temperature Control unit LED reading	______

Calibrated temperature meter reading 	______

X.	Object Alignment

#### Description									 	 pass, fail, N/A

## Locate the object reference position of the Tray-holders or Stacks by locating 		______        the hole in the center of the Tray-holders or Stacks (top drawer must be opened          Manually to access).

- Adjust the X, Y and Z positions and verify by using the F1 Function key 		______ “Check Positions” for each Tray-holder or Stack.

- Locate the Wash Station’s X, Y and Z positions for Wash1 and Wash2. Verify by 	______     using the F1 Function key  “Check Positions”.

- Use the Change Syringe Function, insert the syringe into the syringe holder on the    	______ Injection Head.

- Prime the Syringe utilizing the “Clean Syr” with the “Wash1 and/or the Wash2”       	______ Solvents.

- Define the Tray sizes to be used with the Tray-holders or Stacks.			______

- Locate the Tray’s  X, Y and Z positions for Sample positions. Verify by using 		______        the F1 Function key  “Check Positions”. Repeat for each Tray.

- Locate all inlets X, Y and Z positions. Verify by using F1 key and penetrate the 	______    septa by checking the 40 mm injection depth.

- Verify all other syringe types if applicable. Verify the proper detection of each           	______ syringe.

- Attach the Remote Cable to the PAL Prep and Load System and/or the LC and/or 	______ the Pump and/or the Mass Spectrometer.

- Define the Sync Signals and the injected Output Signals.				______

- Attach the Serial cable provided to customer PC and to the PAL Prep and Load           ______ System. If customer PC unavailable use Portable PC.

- Install the Pal Loader program and Cruise Control appropriate Drivers for 	   	______ Chemstation or Xcalibur (if available) unto the PC.

- Run the PAL Loader program and make a Backup of the PAL system. Keep a 		______      copy for the customer and send another copy to LEAP Technologies.

- Reset the PAL Prep and Load system after completion.

- Disconnect the serial cable from the portable PC.					______

- If the PAL Prep and Load System is controlled by the Mass Spec software, leave 	______ the Serial cable attached to the PC of Mass Spec and the PAL Prep and Load System.

- The PAL Prep and Load System is ready for use.					______

XI.	Functionality Test

#### Description									 	 pass, fail, N/A

- Add 2mL vials to the tray and check following positions via the handheld remote:	______ Setup\Objects\Trays\Select Tray\F3 and go to Position. A 2 mm distance to the vial is acceptable.

| Tray and Well Plate Position Check |  |  |
| --- | --- | --- |
| Position | Indicate Pass/Fail | Verified by |
| Microtiter Stack (MT96 / MT384) if applicable: |  |  |
| Tray1: 001, 012/024, 096/384 |  |  |
| Tray12: 001, 012/024, 096/384 |  |  |
| Deep-well Stack (VT 54 / DW96) if applicable: |  |  |
| Tray1: 001, 009/012, 054/096 |  |  |
| Tray6: 001, 009/012, 054/096 |  |  |
| Cooled Stack (VT 54 / DW96) if applicable: |  |  |
| Tray1: 001, 009/012, 054/096 |  |  |
| Tray6: 001, 009/012, 054/096 |  |  |

- Change the syringe a check their automatic detection. Use the Menu/F1 function 	______         to exchange the syringe. Make sure to check no syringe as well.

- Verify the sync and start signals. Start a new job and program a short sequence 	______         on the PAL. Use following variables:

| ## HTS Pal conditions |  |
| --- | --- |
| Cycle | Liquid Injection |
| Syringe | 100 uL |
| Fill Speed | 5 uL/s |
| Inject Speed | 100 uL/s |
| Inject to | LC valve |
| Pre Inject Delay | 0 ms |
| Post Inject Delay | 0 ms |
| Pre Clean Syringe w Solvent One | 3 |
| Pre Clean Syringe w Solvent Two | 0 |
| Post Clean Syringe w Solvent One | 0 |
| Post Clean Syringe w Solvent Two | 0 |
| Clean w sample | 2 |

Start a short sequence of at least 2 samples to verify communication between the Pal and the LC.

- Check the Pal Loader Software for functionality and take a backup of the current    	______ firmware.

- If purchased, check the Cruise Control or Cycle Composer for functionality. Run a 	______ short sequence with Cycle Composer and make sure that the system starts the external systems. Reload the Pal Configuration and make the Customer aware of the option 	 Format Sample List under File/Options/Format Sample List. Depending on which 	 macros are activated different objects in the sample list are necessary.

- Training / Familiarization

This is essential for a new installation of the PAL autosampler. Only trained users should operate the PAL autosampler. Usually, several users will work more or less frequently with the PAL. It is good if one user will take over responsibility for any questions regarding maintenance, manuals, and technical expertise.

The training should focus on usability, basic method development, routine maintenance, basic programming via the handheld remote, starting signal configuration, realigning the system and backup of the system. Emphasize the importance of a regular backup and routine maintenance.

___________________	      __________________

Attendee			Phone number

___________________	      __________________

Attendee			Phone number

___________________	      __________________

Attendee			Phone number

___________________	      __________________

Attendee			Phone number

XIII.	Enclosures

The following documents are filed and stored in the correspondence LOG Book for the PAL System.

Copy of P.O.

Delivery Note

Certificate of Test Conformation

Non-Conformance Record

Certification letter of service engineer

Certificate of Equipment used for validation (Temp Probe and meter: and Balance		      Tape Read-outs -if available)

Other documentation	 __________________

## Certification

Instrument Qualification

# CERTIFICATE OF CONFORMANCE

## HTS-PAL

Serial Number: ________________

#### Has successfully passed all installation qualifying specifications.

Date________________________

________________________________________________	______________

### Certification Engineer									Date

________________________________________________	______________

## Customer Acceptance 								Date

## Non-Conformance Record

| Ref # | Detailed Description of Non- Conformance | Action taken | Accepted changes Initial and Date by both parties |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| 4 |  |  |  |
| 5 |  |  |  |
| 6 |  |  |  |
| 7 |  |  |  |
| 8 |  |  |  |

# LEAP HTS-PAL IQ Validation Document

Execution Final Approval

The signatures below represent approval of all interested parties for the final acceptance of the completed LEAP HTS-PAL Installation Qualification Protocol.

Company Name____________________________________________________:

Name:			Signature:			Title:			Date:

## LEAP Technologies Representative

Name:			Signature:			Title:			Date: